Litz Wire Termination

- Switch on the solder bath heater

- Wait for it to reach 330°. In the meanwhile, cut your termination to leave around 3-4cm (for test) or 1-3cm (for final soldering).

- Take the transformer in your hand without touching the wires

- Dip the wire in the tin and wait 15-25 seconds, until the isolation layer burns

- Wait for the wire to cool down and remove the burned residuals

- WARNING: Don’t touch the wire with your ends it’s going to be hot!

# LCR Meter and Winding Direction Test

MEASUREMENT OF THE MAGNETIZING INDUCTANCE

- Set the Ls-Rs mode (Press the arrows to select FUNC, select select Ls-, then select Ls-Rs).

- Input your resonance frequency. (Press the arrows to select FREQ, then insert the desired frequency, e.g., 150 with the number pad, then press kHz).

- Connect the black and red clamps to your primary winding.

- Keep the other windings open.

- Measure your magnetizing inductance (value on the screen).

MEASUREMENT OF THE LEAKAGE INDUCTANCE

- Set the Ls-Rs mode (same as before).

- Input your operating frequency (same as before).

- Short circuit your secondary winding.

- Connect the black and red clamps to your primary winding.

- Measure your leakage inductance (value on the screen) with the secondary side in short circuit.

- Measure your series resistance (value on the screen).

SWITCH OFF THE DEVICEs IF NO ONE IS USING IT AFTER YOU

# Bode 100 Test

- Plug Bode 100

- Open Software: Bode Analyzer Suite

  - Select Impedance Analysis

  - One-Port -> Select measurement

  - ONLY NEEDED IF YOU ARE THE FIRST TO PERFORM THE TEST: Full Range -> Perform new calibration -> Open (test point open) Start -> Short (test point on the SHORT part of the Calibration Device) Start -> Load (test point on the LOAD part of the Calibration Device + Advanced Settings 100Ω) Start

- Cursor 1 -> Frequency -> Set manually to your operating frequency

- Trace 1 -> Format -> Ls

- Trace 2 -> Format -> Rs

- Connect the test points to your primary (Secondary open or short circuited according to the test you want to perform)

- Press Single Measurement

SWITCH OFF THE DEVICE IF NO ONE IS USING IT AFTER YOU

# Power Check Test

- Turn on the Test device (Botton on the back)

- Open Software: DPG10 Power Chock Test

  - Max current -> Peak current on your primary

  - Max voltage -> 15 V

  - Pulse time max ->20 us (if it is too low the graph will not reach the maximum current)

- Check that “Voltage Sense” is connected.

- Check that  “1A … 10A” is connected.

- Connect the black and rad clamps to your primary

- Cover inductor with Plexiglass box.

- DO NOT TOUCH ANYMORE

- Run experiment (blu arrow).

- Wait until experiment is finished.

- Save results on USB.

SWITCH OFF THE DEVICE IF NO ONE IS USING IT AFTER YOU

# Windings Heat-Run Test

- Check with a TA for assistance

# Core Excitation Test

- Check with a TA for assistance